TECHNICAL SPECIFICATION OF DELIVERY

Hall measurement system

Device description: "The subject of the delivery is the Hall measurement system. The system must be useful for the determination of following quantities needed for detailed description of materials electrical behaviour: electrical resistivity, concentration and mobility of free carriers in semiconductors.”

- Hall measurement system

- Resistivity measurement range min. from 10-04  to 109  Ω-cm

- Mobility measurement from 1 to 107 V/s

- Measurement of carrier concentration in range: 107 to 1021 per cm3

- Minimum hall voltage measurable: 10 µV

- Magnetic field: minimum 1 Tesla permanent magnet, magnetic field diameter minimum 70 mm

- Automated movement of magnets

- Samples with a minimum size of 6 x 6 mm up to a maximum size of 12 x 12 mm with a maximum thickness of 2 mm can be measured on the device.

- Fully automated measurement sequence

- Standard samples with printed test data

- Laptop with Windows 10 and Hall measurement software installed. All necessary software controllers and drivers are provided on CD or other digital media (USB flash disk, etc.)

- The measurement system software capabilities: Enables data logging and plotting of different measured quantities, I-V curve, resistance, resisitivity, sheet resistance, magnetoresistance, carrier concentration, Hall mobility, Hall coefficient … all as a function of sample temperature.

- Measurement Head – cooling and heating

- Temperature range at least from 100K to 800K with a resolution of ± 0.2 K.

Cooling is possible by liquid nitrogen (LN2).

- Heater controlled by PID temperature controller in the temperature range from 300 to 800 K.

- Entire temperature range in a single system with replaceable head for low (100 – 300K) and high (300-800K) temperature measurements.

- Accessories

- Dry vacuum pump including all necessary pumping lines from pump to chamber with the measured sample, clamps, o-rings, connecting adaptors, vacuum grease, spare   o-ring for vacuum lid.

- Soldering iron, silver paint, tweezer – for making contacts.

- The device contain Dewar container for amount 2ltr for liquid nitrogen (LN2). In the delivery must be Teflon funnel.

- Other terms of delivery

- Complete installation and training at least for 3 persons. The training could be guided on-line communication tool (Skype, Zoom, etc.)

- User manual in English eventual in Czech languages.

- Installation at least in December 2020.

- Warranty valid at least for a period of 24 month.

- Packing and handling are be included in the final price.